## Slide 1: Bandwidths of Transmission Line Anodes
Jean-Francois GENAT
LPNHE Paris
LAPPD Workshop, Chicago, April 26th 2011

## Slide 2: LAPPD 8” x 8” Micro-channel Plates
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Micro-strip Lines
Parameters:
- Strip width W
- Inter-strip distance d
- Metal conductivity s
- Dielectric constant e r

## Slide 3: Timing resolution
(Stefan Ritt)
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Since
and
where
whith

## Slide 4: MCP signals
Segmented anodes, Burle-Photonis Glass + ALD MCPs (from Matt Wetstein, ANL)
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Rise-time 66ps
abw=5.3 GHz

## Slide 5: MCP ‘nominal’ signals spectra
25 microns at 2 kV, 50 Photo-electrons
10 microns at 2.5 kV “
Data taken at Argonne
(Ed May’s laser test stand)
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011

## Slide 6: ‘Nominal’ Timing resolution (model)
Assume:
SN = 50
fs= 10 GS/s
abw = 1 GHz
Tr = .35/abw = 350ps
Ts = 1/ fs = 100ps
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011

## Slide 7: Skin effect
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
No need to have more than 10mm metal (~5 x d at 1 GHz)
Resistivity and
Skin Depth at 1 GHz
Al 2.8 nW.m 2.4 mm
Au 2.4 2.4
Cu 1.7 2.0
Ag 1.6 1.9

## Slide 8: Test plate
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Notes: 8

## Slide 9: Measured Microstrip Bandwidth
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
3dB bandwidth is 3.77GHz for a 8” trace
This measured bandwidth agrees with the simulations using HFSS (Mentor-Graphics)
Notes: 9

## Slide 10: Transmission Lines Bandwidths
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Two main factors for losses:
- Resistance of the metal traces vs 50 W: metal conductivity, skin effect, as
- Dielectric losses: complex dielectric constant
The Borofloat glass loss tangent is unknown at 10 GHz, but we measured it at 3.77 GHz for
a 8 ’’ long microstrip line.
The loss tangent at 10 GHz can be deduced from this measurement to
Copper
10 GHz
A 5m copper line has a metal induced 3dB bandwidth of 10 GHz

## Slide 11: Examples with Quartz and Alumina
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Dielectric losses
Example: f=10GHz, tan(d)= 6 10-5 , C=60pF/m, Z0=50W
Quartz
Alumina
The strip length for a 3dB bandwidth is:
At 10 GHz:
Quartz: 58m
Alumina: 35m
Notes: 11

## Slide 12
Thanks !

## Slide 13: Borofloat 33 properties
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Notes: 13

## Slide 14: Dielectric constants
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Quartz 3.8 1.15 10-5
Beryllia 6.7 1.4 10-5
Fused Quartz 3.8 5 10-5
Alumina 96% 10 2.0 10-4
Alumina 99.5% 9.6 3.0 10-4 10 GHz
Glass (Borofloat) 4.6 3.7 10-3 1 MHz 5.0 10-3 3.7GHz
RF ceramic 6.1 3 10-3
Dielectric constant Loss tangent @ 1 GHz
Notes: 14

## Slide 15: Example: FR4 Microstrip line Losses
Jean-Francois Genat, LAPPD Workshop, Chicago, April 28th 2011
Dielectric FR4 losses dominate > 200 MHz
Conductor losses contribute > 5 GHz
Andrew Byers (Tektronix)
Notes: 15